## Slide 1: 4K Magnet Cryogenic System Problem Solving
MBG Latham, NY
March 30, 2013

## Slide 2: 4K Magnet System Problems
4K magnet system problems can be identified in several ways:
  Difficulty in fully inserting the high current leads
  Difficulty in achieving proper lead voltage drop
  Poor helium fill transfer efficiency
  Low magnet margin
  Loss of magnet pressure control
Items 1, 2, & 3 are indications that there could be iced particle contamination in the high current lead guides, on the high current lead terminals or on the terminal board. This iced particle contamination must be removed by de-icing.
Items 4 & 5 may also be indicators of iced particle contamination but, more commonly are caused by problems with the cryogenic support system.

## Slide 3: Low Magnet Margin
Is this a new or existing installation?
  If this is a new installation, check for proper installation of system
  If this is an existing installation, review MEU cryogenic logs for faults and correct any faults discovered
Has the magnet margin ever been correct?
  How much margin and for how long?
Does the MEU have uninterrupted power?
Does the coldhead run continuously?
  Review short term cryogenic log for interruptions
Is the helium compressor cooling water flow correct?
  Water flow should normally be 8 L/min @ 10°C
Is the helium compressor charge pressure correct?
  Static charge pressure should be 1.65 – 1.70 Mpa
  Dynamic operating pressure should be 2.1 – 2.4 MPa

## Slide 4: Low Magnet Margin
How long has it been since the last coldhead service?
  Less than 36 months; perform 4K Coldhead Gas Purification Procedure
    Does margin improve?
      No; coldhead may require service
      Yes; does it return to normal?
        Margin returns to normal; procedure complete
        Margin is still low; perform 4K Coldhead Gas Purification Procedure again
          If performance does not improve further, coldhead service may be required.
  Longer than 36 months; coldhead service may be required

## Slide 5: Loss of Magnet Pressure Control
Is this a new or existing installation?
  If this is a new installation, check for proper installation of system
  If this is an existing installation, review MEU cryogenic logs for faults and correct any faults discovered
Has the magnet ever been in pressure control?
  Review short term and long term cryogenic logs
    Magnet was in pressure control at some point
      What changed?
Does the coldhead run continuously?
  Review short term cryogenic log for interruptions
Is the helium compressor bypass light “ON”

## Slide 6: Loss of Magnet Pressure Control
Is the helium compressor cooling water flow correct?
  Water flow should normally be 8 L/min @ 10°C
Is the helium compressor gas charge pressure correct?
  Static charge pressure should be 1.65 – 1.70 Mpa
  Dynamic operating pressure should be 2.1 – 2.4 Mpa
Verify modified wet-sock communication with 4K vessel
  Open one port on the modified wet-sock and one port on the main turret
    Flow should be equal and steady from both ports for at least 30 seconds
    If there is any difference between flows, or if flow decreases from one of the ports, perform 4K Magnet De-ice Procedure

## Slide 7: Loss of Magnet Pressure Control
Is magnet pressure stable at less than 240 hPa?
  How long has it been since the last coldhead service?
    Less than 36 months; perform 4K Coldhead Gas Purification Procedure
      Does margin improve?
        No; coldhead may require service
        Yes; does it return to normal?
          Margin returns to normal; procedure complete
          Margin is still low; perform 4K Coldhead Gas Purification Procedure again
            If performance does not improve further, coldhead service may be required.
  Longer than 36 months; coldhead service may be required

## Slide 8: Loss of Magnet Pressure Control
Is magnet pressure 240 hPa or higher?
  Perform 4K Magnet De-icing Procedure
  Monitor magnet for pressure control and margin after completion of the 4K Magnet De-icing Procedure