PRIVATE
TP120 - Temperature & Humidity Data Logger with 5 Year Battery - $159 

Category: Temperature and Humidity 
Product type: PC Based Products 

Order Now | New Search

Features: The Dickson Pro Series Data Loggers take data acquisition to the next level with 5-year battery life, data storage capacity of 32,512 samples and our best accuracy specifications. 

The TP120 is a self-contained, temperture & humidity Data Logger equipped with internal sensors for recording temperature and humidity. 
User friendly DicksonWare software creates a graph of your data in seconds. DicksonWare features include complete customization of your graph, zooming, and viewing your data in tabular format. You can also view data from several loggers on a single graph enabling easy comparison of multiple sets of data. Transferring the data into documents or spreadsheets is a snap! 


· Battery Operation - 5 year battery life 

· Stores 32,512 Sample Points 

· Temperature Range -40 to +176°F (-40 to +80°C) 

· Humidity Range 0 to 95% RH (non-condensing) 

· "Unit On" LED Indicator 

· 2% RH Accuracy 

· 30 Day Money Back Guarantee 



showpic.cgi?product_id=81

INCLUDEPICTURE  \d  \z "product_images/TP120.gif"


HYPERLINK "showpic.cgi?product_id=81"
showpic.cgi?product_id=81

Click photo to enlarge 

Detailed Specs 
Models 
Accessories 
Charts 
Technical Support 


Specifications (summary): 

Temperature Range: -40 to +176°F (-40 to +80°C)Humidity Range: 0 - 95% RH (non-condensing) 
Power Source: Lithium Battery 5 Year (Factory Replaceable) 
Software: Microsoft 95, 98, 2000 and NT compatible 
Data Capacity: Single Channel - 32,512 sample points Dual Channel - 16,256 sample points per Channel 






..\contact.htmlPRIVATE "TYPE=PICT;ALT=Contact Information"


HYPERLINK "../contact.html"
..\contact.html
PRIVATE
TP120 - Temperature & Humidity Data Logger with 5 Year Battery - $159 

Category: Temperature and Humidity 
Product type: PC Based Products 

Order Now | New Search 
TP120 Product Information 
Detailed Specs 
Models 
Accessories 
Charts 
Technical Support 


Detailed Specifications 

Temperature Range: -40 to +176°F, -40 to +80°C Temperature Sensor: Digital temperature sensor 
Temperature Accuracy:>±1.8°F(±1°C) over range +14 to +176°F (-10 to +80°C); ±3.6°F(±2°C) over range +14 to -40°F (-10 to -40°C) 
Humidity Range: 0 - 95% RH (non-condensing) 
Humidity Sensor: Monolithic Humidity Sensor 
Humidity Accuracy: ±2% at 40% RH and +73°F (+23°C) 
±4% over range 0 to 95% RH at +41 to +122°F (5 to +50°C) 
Resolution:Internal Temperature Sensor: 0.1°F (0.1°C) Internal Humidity Sensor: 0.1% RH 
Power Source: Lithium CR-2477 3V (Factory Replaceable) 
Battery Life: 5 Years typical @ 1 minute sample intervals 
PC Software Features:Windows™ based, variable sample interval, data wrap-around, viewing data, graphing and zooming 
Data Capacity: Single Channel - 32,512 sample points Dual Channel - 16,256 sample points per Channel 
Sample Interval:Variable from 10 seconds to 24 hours in 10 sec increments 
Response Time:10 minutes to move 63% of full scale (temperature) and 15 minutes to move 63% of full scale (humidity) 
Download Time: 3 minutes, typical with logger full 
Power Status: LED - Flashes when full 
Ambient Operating Range: -40 to +176°F (-40 to +80°C) 0 -95% RH (non-condensing) 
Approvals:CE 
Dimensions: 3.13" x 2.14" x 0.87" (79mm x 54mm x 22mm) 
Memory: Non-volatile 
Calibration: User: Zero Pt w/SW400 Software; Factory: Zero Pt. and Span 
Enclosure: Polycarbonate 
Weight: 1.8 oz (49g) 

A015 - DicksonWare & Download Cable Specifications 
Software: Microsoft 95, 98, 2000 and NT compatible 
PC Requirements: PC w/386 or better micro processor, 4M RAM, 1 free COM (Serial) port, CD drive 
Version Required (Minimum): 5.0.14 
Cable Type/Length: 9 pin male D-shell to male 2.5mm sterio plug, 6 feet long 
Computer Interface: RS-232 COM (Serial) port 

Note: A015 must accompany the initial purchase of the TP120 in order to operate. 


..\contact.htmlPRIVATE "TYPE=PICT;ALT=Contact Information"


HYPERLINK "../contact.html"
..\contact.html
PRIVATE
TP120 - Temperature & Humidity Data Logger with 5 Year Battery - $159 

Category: Temperature and Humidity 
Product type: PC Based Products 

New Search 
TP120 Product Information 
Detailed Specs 
Models 
Accessories 
Charts 
Technical Support 


Models 


Top of Form 1




PRIVATE
Model
Description
Price
Qty.

TP120
Temperature & Humidity Data Logger with 5 Year Battery
$159
[image: image1.wmf]




Notes: A015 must accompany the initial purchase of data logger in order to operate 




A015
DicksonWare Software and Download Cable
$59
[image: image2.wmf]




Notes: Required with initial purchase of any Dickson Logger



[image: image3.wmf]

Add to Order


Bottom of Form 1




..\contact.htmlPRIVATE "TYPE=PICT;ALT=Contact Information"


HYPERLINK "../contact.html"
..\contact.html
PRIVATE
TP120 - Temperature & Humidity Data Logger with 5 Year Battery - $159 

Category: Temperature and Humidity 
Product type: PC Based Products 

New Search 
TP120 Product Information 
Detailed Specs 
Models 
Accessories 
Charts 
Technical Support 


Accessories 


Top of Form 1




PRIVATE
Part Number
Description
Price
Qty. 

A015
Software & Interface Cable
$59
[image: image4.wmf]



A713
Water Resistant Pouch
$25
[image: image5.wmf]



A715
Locking Wall Mount Case
$40
[image: image6.wmf]



E100
One Year Extended Warranty 

$59
[image: image7.wmf]



N100
1-Point NIST Traceable Calibration
$99
[image: image8.wmf]



N300
3-Point NIST Traceable Calibration
$149
[image: image9.wmf]



N400
A2LA Accredited Calibration 3-pt.
$259
[image: image10.wmf]



N520
Software Certificate of Validation
$29
[image: image11.wmf]



SW400
Calibration Software
$149
[image: image12.wmf]



[image: image13.wmf]

Add to Order


Bottom of Form 1




..\contact.htmlPRIVATE "TYPE=PICT;ALT=Contact Information"


HYPERLINK "../contact.html"
..\contact.html
_1043856571.unknown

_1043856574.unknown

_1043856575.unknown

_1043856572.unknown

_1043856565.unknown

_1043856568.unknown

_1043856569.unknown

_1043856566.unknown

_1043856495.unknown

_1043856562.unknown

_1043856563.unknown

_1043856497.unknown

_1043856485.unknown

