| Philips Healthcare |  |  |  |  |  |  |  |
| --- | --- | --- | --- | --- | --- | --- | --- |
| NEWSLETTER |  |  |  |  |  |  | MR Service |
| Issued by | : | Philips Healthcare; Best |  | Reference No. |  | : | MR-223 |
| Publication nr | : | DMR113812 |  | Date |  | : | March 2010 |
| Information for | : | FSE | Managers |  | Specialists |  |  |
| Extended cold head rebuild: |  |  |  |  |  |  |  |

# Introduction

Up to now the rebuild of the cold head (part of the refrigerator system) is being rebuild every 2 years. After years of experience we introduce with this newsletter a new rebuild interval which is an optimum between:

- costs of cold head parts

- cost of liquid helium

- risks of corrective actions

## Conclusion

The new rebuild interval is depending on how the performance of the refrigerator system is being watched. The longer a cold head is in the system without a rebuild the higher the risk of a dip in the performance or complete shutdown. If the refrigerator system is checked on a regular basis the rebuild interval is longer.

New recommended service intervals:

| Cold head type | Planned rebuild interval | Maximum interval with regular performance check |
| --- | --- | --- |
| 10K cold head | 3 year | 4 year |
| 4K cold head | 3 year | 4 year |
| adsorber | 3 year | 4 year |

Contents

1	Introduction	1

1.1	Conclusion	1

1.2	Scope	3

1.3	Legend	3

1.4	Part numbers	3

1.5	Parameters	3

1.6	Cold head PM part of the new (10K) PM manual	4

1.6.1	Check of the cryogenic performance of the (10K) magnet system	4

1.7	Cold head PM part of the new (4K) PM manual	4

1.7.1	Check of the cryogenic performance of the (4K) magnet system	4

1.7.2	Determining the moment of coldhead rebuild	5

## Scope

This newsletter is valid for the following magnet systems:

F2000 magnets (10K)

HFO magnet systems (4K)

Titan magnets 3T (10K)

Rex Magnets 3T (4K)

## Legend

| Legend | Description |
| --- | --- |
| Cold head | Part of the refrigerator system mounted on top of the magnet |
| Displacer | Part of the cold head which is moving up and down |
| Cylinder | Part of the cold head which is welded inside the magnet |
| Compressor | Part of the refrigerator system located in the technical room |
| Refrigerator system | Combination of the compressor gas lines and cold head |
| F2000 | Cylindrical magnet 0,5T 1T and1.5T Al F2000 magnet have a 10K cold head |
| HFO | High Field Open magnet; Open 1T magnet used in Panorama systems, always 4K |
| Titan | First Philips 3T magnets with a look of an F2000 magnet, Always 10K |
| Rex | 2e version of 3T magnets. Always 4K |
| 10K | Temperature of the cold head in F2000 magnets, Reduce the boil off |
| 4K | Temperature of the cold head in HFO and REX magnets. |

## Part numbers

| Refrigerator part | 12nc |
| --- | --- |
| Compressor HC8-C4 | 4522 150 12241 |
| Compressor HC8-E | 4522 150 22982 |
| Compressor F50 | 4522 133 06131 |
| Cold head SC-8C | 4522 150 11821 |
| Cold head F2000 | 4522 150 22732 |
| Replacement parts 408C2 | 4522 150 42991 |
| Replacement parts 415 C2 | 4522 150 37731 |
| 10k refrigerator maintenance tool | 4522 150 21652 |
| 4k refrigerator maintenance tool | 4522 150 42711 |
| Adsorber 4K | 4522 150 38241 |
| Adsorber 10K | 4522 130 74033 |

## Parameters

Extending the cold head rebuild will reduce the cost of planned maintenance. Extending the maintenance interval has also a risk. A not performing cold head or a cold head which is down will increase the boil off. In some situations the displacer of the cold head will break. If a displacer is broken a 3e line action will have to be performed with use of special tools to get the broken part out of the magnet.

In extreme situations the complete refrigerator system will have to be exchanged because of contaminated helium gas. Although we have never run in this problem there is a risk that the cylinder of the magnet will be damaged by the displacer. If a cylinder of a F2000 magnet is damaged the magnet has to be returned to the factory for repair

Taken all these parameters into account results in recommended rebuild intervals. See the table on page 1.

## Cold head PM part of the new (10K) PM manual

Perform a planned rebuild every 3 year. If the performance of the refrigerator system is still ok it is allowed to extend the rebuild to maximum 4 years. See next paragraph to check the cryogenic performance of a 10K magnet system.

### Check of the cryogenic performance of the (10K) magnet system

To check the helium boil-off, you have to use the daily table of the helium level, entered by the customer or from release R6.1, logged by the software. Calculate the boil-off over a long period (longer than 3 weeks). Look in the logbook to make sure that there were no cryogenic actions or magnetic field changes in the chosen period. Calculate the boil-off as follows:

The capacity of the magnets (100% full) is and maximum boil off is:

| F2000 magnet types | 1650 litre maximum boil off 0,075 litre/hour |
| --- | --- |
| 3T magnet | 1500 litre maximum boil off 0,200 litre/hour |

## Cold head PM part of the new (4K) PM manual

Perform a planned rebuild every 3 year. If the performance of the refrigerator system is still ok is allowed to extend the rebuild to maximum 4 years. See next paragraph to check the cryogenic performance of a 4K magnet system.

### Check of the cryogenic performance of the (4K) magnet system

|  | Make a graphs (*), with on the horizontal axis time and on the vertical axis: - The helium level in % or for 1T panorama magnets: - The helium level in % in the upper bath. - The helium level in % in the lower bath. If the helium levels decrease, there is a leak. Find the leak and repair it. In case of a leak continue with step 2 after a month. |
| --- | --- |
|  | Make a graph(*) , with on the horizontal axis time and on the vertical axis: - The power to the bath heater. - The helium pressure in the cryostat. |
|  | Conclude if the refrigerator system degrades: - Has the cold head been operating for more than 3 years? - Does the bath heater power show a bath curve similar to Figure 1 over at least a year? - Has the helium pressure risen above 150 hPa? If the answer to the 3 questions is “Yes”: Or the coldhead is operational > 4 years plan a cold head rebuild. If the cryogenic performance is not ok and the cold head operated < 3 years, perform corrective actions like: - Setting the pressure of the compressor - Correct the water (temp / flow) of the compressor - De-icing - Charge run and Vent procedure (*) tools and description to make a graphs using the STT_magne.txt file, will be available on InCenter. |

### Determining the moment of coldhead rebuild

Figure 1 - Typical graph of degradation of the cold head

4 years

(approx.))

Power to bath heater

Helium pressure

150 hPa pressure

30 hPa pressure

years

In the figure above, the power to the bath heater decreases slowly in time. The cause is that the cold head cooling power decreases. If the cooling power is reduced, the boil-off simply increases and the helium pressure rises. At a helium pressure of 150 hPa mBar the cold head has to be replaced.