Signa Ovation RF Coil Feedback Sheet
The coils listed in data work sheet are the M3 coils (pilot unit). We need your feedback to improve the quality. If you install the MFO coil, please fill this sheet out and send this back to the service engineering representatives in each region. 

Installation information
Comment

FE name



Hospital name



Installation date(yyyy/MM/DD)



Total Installation time


Coil numbers
coils


Tuning time (if required.)
minutes
             

Scan time
minutes
             

Others
minutes
             

SNR Data Work Sheet

Coil
Ch.
Spec.
SNR
Coil
Ch.
Spec.
SNR

CTL-C
Composite
280.40

Shoulder
Composite
155


CTL-T(P50,S100)
Composite
93.48


2
n/a


CTL-T(P50,S0)
Composite
109.32


4
n/a


CTL-T(P50,I100)
Composite
125.16

Wrist
Composite
324


CTL-L(P50,S100)
Composite
99.89


1
n/a


CTL-L(P50,I20)
Composite
122.36


2
n/a


CTL-L(P50,I100)
Composite
103.91


3
n/a


Knee Large
Composite
175

Breast-R
Composite
113



5
n/a


1
n/a



7
n/a


2
n/a






Breast-L
Composite
113







3
n/a







4
n/a


Service engineering representative; 

(email is also available. You can obtain this sheet from Service CD-ROM rev. 3 or later.)


Americas : Rob Ash (robert.ash@med.ge.com)



3200 N Grandview Blvd, Waukesha, WI 53188, USA


Asia, Europe : Haruhiko Akeda (haruhiko.akeda@gemsa.med.ge.com)



4-7-127, Asahigaoka, Hino-shi, Tokyo, 191-8503 Japan

2001.12.20

